Design and optimization of patient-specific, pediatric laryngoscopes

###### R. Sims1, M. Boutelle1, J. Inziello1, and F. Lobo1\*

1 PD3D Lab, University of Central Florida, Orlando, United States of America

\* Corresponding author, email: [jstubbs@ist.ucf.edu](mailto:jstubbs@ist.ucf.edu)

Abstract: Place a brief summary of your work here. Do not use more than 100 words. 3D printing is of outstanding importance in medical engineering and has been growing continuously in recent years. From prostheses and soft implants to matrices for tissue engineering, additive manufacturing has decisive advantages for medicine. The scientific conference AMMM 2019 brings together engineers, scientists and technicians with physicians and entrepreneurs to discuss the latest achievements in 3D printing development for medicine.

# I. Introduction

Originally developed for otolaryngologists to inspect vocal cords, laryngoscopes have undergone continuous modifications since their inception, eventually finding a place in anesthesiology. In 1911, Dr. Chevalier Jackson developed the first laryngoscope that allowed for the insertion of an endotracheal tube (ETT) [1]. In the same year, Dr. Henry Janeway introduced a battery powered, distal light source allowing for optimized viewing conditions [1]. Modern laryngoscopes, such as the Macintosh and Miller, began manufacturing in the early 1940’s. The Macintosh’s continuous curved blade allots more room in the oropharynx for successful passage of the ETT. The Miller’s straight blade design, with curved distal tip, provides an improved view of the glottis [2]. In the last few decades, laryngoscope design changes have focused on addressing challenging airways. Most modern laryngoscopes, such as the McGrath, Glidescope and Airtraq, feature integrated optics and video screens. Additionally, the three brands also feature variable-size, single-use (disposable) blades. Blade sizes are distributed unevenly across adults (3-4 sizes), pediatrics (one size), and neonates (one size). Sizes match a range of ETT sizes (2.5-3.5 for neonate, and 4.0-5.5 for pediatrics) [3].

As the industry moves in the direction of single-use medical devices, there is potential to shift from size groups to patient-specific blades. This is relevant to pediatric and neonatal cases, were size options are limited. Difficulties with intubation represent the main cause of pediatric, anesthesia-related morbidities and mortality [4]. Even in scenarios where difficult intubations are expected, anesthesiologists know to have “all the equipment to hand” [5]. Patient-specific blades would ensure readiness in the case of normal and abnormal airways.

The development of patient specific devices requires the integration of advanced reconstruction, design, and manufacturing technologies. Using SideFX’s Houdini, we have consolidated the design process into a single program (Figure 1).

Figure 1: Houdini Program Flowchart

# II. Material and methods

## II.I. Patient Data

De-identified or anonymized CT DICOM datasets from three (3) pediatric cases were obtained from Nemours Children’s Hospital (Lake Nona, FL, USA). In addition to age, the only additional information retrieved were notes solely related to findings that could have an effect on the patient’s airway.

The first patient, 18 months old, did not present any lesions or abnormalities affecting the airway. The second patient, 2 months old, presented a well-defined ovoid lesion measuring approximately 3.0 cm x 3.6 cm x 2.8 cm. The lesion was situated deep to the left lobe of thyroid gland and medial to the left common carotid artery and internal jugular vein. It extends medially in the prevertebral space across the midline; anteriorly the lesion displaces the left lobe of the thyroid. The resulting airway is deviated to the right side from the effect of the lesion. The third patient, also 18 months old, presented a lesion on the right side of the neck. The lesion compromised the patient’s oral cavity, oropharynx, and nasopharynx, resulting in the slight narrowing of the airway.

## II.II. Segmentation

Relying on Houdini’s Python 2.x compatibility and Pydicom [ref], our team built custom functions (or nodes in Houdini) to: (1) Import CT DICOM dataset into a voxel (3D pixel) volume, (2) Map the associated Hounsfield data to a [-1, 1] density scale, and (3) Pad boundary voxel data to generate a closed geometry.

## II.II. Shrink-Wrapping

Through a series of voxel-based erosions and dilations, an enlarged copy of the segmented geometry was used extract the patient’s airway through a Boolean subtraction.

## II.III. Pathfinding

To generate a build path that conforms to the patient’s airway, a space colonization algorithm was implemented [ref]. The mouth was selected as the starting point and no branching was allowed. The point cloud derived from the subtracted airway geometry was used to solve for a path between the mouth and the epiglottis of the patient. For all three patients, the algorithm was forced to stop after the path reached the epiglottis.

## II.IV. Part Design and Fabrication

The resulting path was used to extrude a cross-section of a laryngoscope blade. The geometry was exported as a surface file (.OBJ) for further editing and fabrication.

# III. Results and discussion


Table 1: Values used to achieve airway curvature

|  |  |  |  |
| --- | --- | --- | --- |
| **Parameter** | **Unit** | **Mean** | **SD** |
| Segmentation\* | Density | 0.2 | 0.0 |
| Voxel Resolution1 | Millimeters | 0.56 | 0.33 |
| Erosion Amount | Millimeters | 2.4 | N/A |
| Dilation Amount | Millimeters | 8.0 | N/A |
| Point Separation2 | Millimeters | 2.33 | 0.94 |
| Frames3 | Number | 8.66 | 1.88 |

1 Using a fixed threshold value of 0.20 (or 200 HU), a closed geometry


Voxel resolution had to be decreased for some scans in order to maintain enough detail to portray the airway, this change is mainly cosmetic and doesn’t affect other calculations.


2 Similarly to the above note, this number had to be changed in order to endure there were enough points scattered in parts of the volume that were otherwise too small.

3 depending on the length of the airway, more or less steps in the solving step may be required.

Talk about why going through positive anatomy instead of negative Anatomy?!?!?!?

Pathfinding is still very manual





# IV. Conclusions

Although a conclusion may review the main points of the paper, do not replicate the abstract as the conclusion. A conclusion might elaborate on the importance of the work or suggest applications and extensions.

### Acknowledgments

##### The preferred spelling of the word “acknowledgment” in America is without an “e” after the “g”. Avoid the stilted expression, “One of us (R. B. G.) thanks . . .” Instead, try “R. B. G. thanks”.

### Author’s statement

##### Research funding: The author state no funding involved. Conflict of interest: Authors state no conflict of interest. Informed consent: Informed consent has been obtained from all individuals included in this study. Ethical approval: The research related to human use complies with all the relevant national regulations, institutional policies and was performed in accordance with the tenets of the Helsinki Declaration, and has been approved by the authors’ institutional review board or equivalent committee.

#### References

[1] M. Weiss and T. Engelhardt, "Proposal for the management of the unexpected difficult pediatric airway," *Pediatric Anesthesia,* no. 20, pp. 454-464, 2010.

[2] R. W. M. Walker and J. Ellwood, "The Management of difficult intubation in children," *Pediatric Anesthesia,* no. 19, pp. 77-87, 2009.